Radiological Clearance/Decommissioning Proposal

| Report from |  |
| --- | --- |
| Position |  |
| School |  |
| Building and room number |  |
| Description of area |  |
| Date of report |  |
| Purpose of clearance |  |

- History of radiation use in laboratory & survey strategy

| Radioisotopes known (or reasonably suspected) to have been used in this area |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Tritium | Carbon 14 |  | Phosphorus 32 or 33 | Sulphur 35 |  |  | Iodine 125 |
| Other details …… |  |  |  |  |  |  |  |
| Insert details of research groups/types of activities performed |  |  |  |  |  |  |  |
| Map of laboratory indicating areas of radiation usage attached (Tick if yes) |  |  |  |  |  |  |  |
| Map of areas to be monitored for contamination attached (see guidance below) (Tick if yes) |  |  |  |  |  |  |  |
| Confirmation that: - all cupboards/drawers/fridges and freezers have been searched - all radioactive materials have been identified - all radioactive materials have been removed |  |  |  |  |  |  |  |
| Comment if necessary: |  |  |  |  |  |  |  |
| Method of contamination monitoring (surfaces) |  |  |  |  |  |  |  |
| Tritium/Carbon 14 |  | Indirect contamination monitoring using wipe tests and liquid scintillation counting |  |  |  |  |  |
|  |  | Wipes taken using |  |  |  |  |  |
|  |  | Details of liquid scintillation counter and programme to be used |  |  |  |  |  |
|  |  | Proposed action level |  |  |  |  |  |
| Phosphorus 32 or 33 |  | Direct contamination monitoring using |  |  |  |  |  |
|  |  | Proposed action level |  |  |  |  |  |
| Sulphur 35 |  | Direct contamination monitoring using |  |  |  |  |  |
|  |  | Proposed action level |  |  |  |  |  |
| Iodine 125 |  | Direct contamination monitoring using |  |  |  |  |  |
|  |  | Proposed action level |  |  |  |  |  |
| Other |  | Direct contamination monitoring using |  |  |  |  |  |
|  |  | Proposed action level |  |  |  |  |  |

| Guidance on monitoring of radiation laboratories |  |
| --- | --- |
| Sinks and draining boards used for the disposal of aqueous radioactive waste | Sinks/drains will be flushed with copious amounts of water. The draining board, sink and plug hole will be monitored. |
| Sink traps from radiation sinks | Following flushing with water the sink traps will be disconnected and monitored internally using a wipe test. |
| Drainage | Open ends of connected pipework will be monitored. Depending on the type of isotopes used and the drainage plans further investigations may be required. |
| Fume hoods used for work with radioactivity | All internal work surfaces and aqueous waste disposal sinks will be monitored as above. Gaseous extract points will be monitored using wipes. |
| Equipment | All equipment in the room will be monitored (all external and internal surfaces) including fridges, freezers and cold rooms. Samples of freezer ice will be counted by liquid scintillation. |
| Benches and cupboards | A matrix of bench space will be monitored, including areas of known usage. Cupboard handles and under bench furniture will be monitored. |
| Floors | A matrix of the floor will be monitored. |
| Further points where radioactive contamination is reasonably foreseeable; e.g. handles, switches etc. | Monitored. |

- Proposed Remediation Strategy

Attempts will made to remove all identified contamination using moist paper wipes initially and then subsequently using [insert method e.g. Decon solution] as appropriate.

Areas where contamination remains fixed in spite of efforts to decontaminate will be highlighted in the decomissioning report to H&S Services.  H&S Services, the School RPS and where necessary the RPA will advise on further action to be taken.